# Protocol: Compendial Method Verification (USP <1226>)

This is a ready-to-use verification protocol for a compendial (pharmacopeial) analytical procedure under USP General Chapter <1226>. Replace every <<FILL: ...>> placeholder with your own specifics, set your document numbers and dates, and get it reviewed and approved before execution. A worked filled specimen follows the template. This is general guidance to adapt, not regulatory advice; confirm the current USP-NF edition and the cited regulation before you rely on them.

# Document control header

| Field | Entry |
| --- | --- |
| Title | Verification of Compendial Procedure: <<FILL: attribute / method>> |
| Document number | <<FILL: protocol ID>> |
| Version | <<FILL>> |
| Status | Pre-approved before execution (required) |
| Author | <<FILL>> |
| Effective date | <<FILL>> |

State clearly in the title: "Verification of Compendial Procedure," not "validation." The wrong title invites the wrong scrutiny.

# 1. Purpose and scope

| Field | Entry |
| --- | --- |
| Article (product / substance) | <<FILL>> |
| Attribute tested | <<FILL: e.g. assay, related substances>> |
| Compendial reference | <<FILL: monograph or general chapter, with the current USP-NF edition and supplement>> |
| Specification the result supports | <<FILL: e.g. 90.0 to 110.0 percent>> |
| Instruments / column lot / CDS | <<FILL>> |
| Analysts in scope | <<FILL>> |

# 2. Regulatory and compendial basis

USP General Chapter <1226>, Verification of Compendial Procedures (cite the current USP-NF edition).

21 CFR 211.194(a)(2): test methods used must be verified under actual conditions of use.

21 CFR 211.165(e): accuracy, sensitivity, specificity, and reproducibility of test methods must be established and documented.

Where chromatographic adjustments are made, USP General Chapter <621> (and Ph. Eur. 2.2.46 equivalent) for allowable adjustments.

Confirm the current chapter content and edition before issue; USP is revised continuously.

# 3. Description of the compendial procedure

Summarise the published procedure as written, and quote the system suitability requirements verbatim so there is no ambiguity about what "passes" means.

| Field | Entry |
| --- | --- |
| Procedure summary | <<FILL: column, mobile phase, detection, sample prep, run conditions as published>> |
| System suitability (verbatim) | <<FILL: resolution, tailing, plate count, replicate %RSD as the monograph states>> |

# 4. Risk assessment summary and characteristics to verify

Reference the characteristic-selection assessment (see assessment-compendial-verification-characteristic-selection). Record which characteristics are verified and the rationale for those excluded.

| Characteristic | Verify? | Rationale |
| --- | --- | --- |
| Specificity / selectivity | <<FILL: usually Yes>> | <<FILL: your matrix is the variable>> |
| Accuracy (recovery) | <<FILL>> | <<FILL>> |
| Repeatability | <<FILL>> | <<FILL>> |
| Intermediate precision | <<FILL>> | <<FILL: yes if more than one analyst/day/instrument>> |
| Detection / quantitation limit | <<FILL>> | <<FILL: low-level impurity work only>> |
| Linearity / range | <<FILL>> | <<FILL: often excluded if published range brackets your spec>> |

# 5. Materials, standards, and instruments

| Item | Detail / traceability |
| --- | --- |
| Reference standard(s) and lot | <<FILL>> |
| Reagents and columns (lot) | <<FILL>> |
| Instruments (ID, qualification status) | <<FILL: confirm AIQ current, see analytical-instrument-qualification>> |
| Software / CDS (version) | <<FILL>> |

# 6. Test design (per selected characteristic)

| Characteristic | Design |
| --- | --- |
| Specificity | <<FILL: inject diluent blank, placebo, placebo spiked with analyte, forced-degradation sample; confirm no interference at the analyte retention time; peak purity if PDA>> |
| Accuracy | <<FILL: spiked placebo at e.g. 80/100/120 percent, triplicate per level, calculate recovery>> |
| Repeatability | <<FILL: six independent preparations at 100 percent, one analyst/day/instrument, report %RSD>> |
| Intermediate precision | <<FILL: second analyst, different day and instrument, repeat preparations; compare>> |

# 7. Acceptance criteria (pre-defined; do not change after seeing data)

| Characteristic | Acceptance criterion (justify per method) |
| --- | --- |
| Specificity | <<FILL: no interfering peak above a stated threshold at the analyte RT; peak purity passes>> |
| Accuracy | <<FILL: mean recovery within a stated band, e.g. 98.0 to 102.0 percent for an assay>> |
| Repeatability | <<FILL: %RSD at or below a method-appropriate limit, e.g. <= 2.0 percent>> |
| Intermediate precision | <<FILL: %RSD within a stated limit; analyst mean difference within a stated absolute>> |
| LOQ/LOD (if tested) | <<FILL: signal-to-noise typically >= 10 (LOQ) and >= 3 (LOD), or the method's stated approach>> |
| System suitability | Meets the compendial SST exactly as written |

# 8. Deviation and OOS handling

Any failure is raised as a deviation and investigated; results are not rerun until they pass. Reference <<FILL: deviation SOP-ID>> and <<FILL: OOS SOP-ID>>. A failed verification is a signal about method suitability for your article, not a reason to quietly adjust the method.

# 9. Roles and approvals

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| Author (QC SME) | <<FILL>> |  |  |
| Reviewer (method validation group) | <<FILL>> |  |  |
| Approver (QA, independent of execution) | <<FILL>> |  |  |

QA approval must be independent of execution; the same person cannot author, execute, and provide the sole quality approval.

──────────────────────────────

# Filled specimen

A pre-approved verification protocol for an example finished-product assay by HPLC, where the formulation has two excipients not in the monograph's typical matrix and two analysts will run it (illustrative).

| Field | Entry |
| --- | --- |
| Compendial reference | USP monograph assay, current USP-NF edition |
| Specification | 90.0 to 110.0 percent of label |
| Characteristics verified | Specificity, accuracy, repeatability, intermediate precision |
| Excluded (with rationale) | Linearity/range (published range brackets the 90 to 110 percent spec; detector linearity well documented); LOQ (major-component assay) |
| Specificity design | Diluent blank, placebo, placebo spiked at nominal, forced-degradation sample |
| Accuracy design | Spiked placebo at 80/100/120 percent, triplicate each |
| Acceptance: accuracy | Mean recovery 98.0 to 102.0 percent per level |
| Acceptance: repeatability | %RSD <= 2.0 percent (n=6) |
| Acceptance: intermediate precision | Combined %RSD (n=12) <= 3.0 percent; analyst mean difference <= 2.0 percent absolute |

The protocol targets the two characteristics the risk assessment flagged (specificity and accuracy, because the matrix is the variable) plus the precision the multi-analyst use demands. It does not redo the full ICH Q2 battery, which is the shape of a <1226> study.

# Common findings this protocol prevents

A compendial method used for release with no verification (a direct 21 CFR 211.194(a)(2) citation).

Acceptance criteria set or moved after the data (testing into compliance).

A modified method called "verified" when the change went beyond <621> allowable adjustments.

No documented rationale for the characteristics that were excluded.

Verification on an unqualified instrument or against a stale compendial edition.

# How to adapt this protocol

Title it precisely and number it under document control.

Quote the actual monograph system suitability verbatim in section 3.

Let the characteristic-selection risk assessment drive section 4.

Justify each acceptance criterion against your spec and matrix, not a copied table.

Get QA pre-approval before any data is generated, and confirm the USP edition is current.